## Slide 1: DOT HRS – Detector Mount

## Slide 2: Scope of this analysis
Estimate heat loads and temps on detector
Estimate XY movement of detector face from heat loads
Compare scenarios below
Estimate thermal transients from detector readout
Scenario | Ribs | Mount | Shield
1 | Yes | As-designed | As-designed
2 | No | As-designed | As-designed
3 | Yes | Ideal flexure mount | As-designed
4 | Yes | Ideal flexure mount | Extended forward

## Slide 3: Mount Overview

## Slide 4: Passive Shield

## Slide 5: Heat loads
Analytically: modelled as cylindrical shell
1 Window
Black Body
e = 1
T = 300K
F12 = 0.06
3 Cryostat Wall
T = 300 K
e = 0.05
F32 = 0.94
2 Detector
(CF @ 160 K)
4 Cold Shield
e = 1
F = 1

## Slide 6: Enclosure 1
1 Window
Black Body
e = 1
T = 300K
3 Cryostat Wall
T = 300 K
e = 0.05
2 Detector
(CF @ 160 K)
4 Shield Face
All view factors from 2 and 4 sum to 1

## Slide 7: Enclosure 2 and 3
Enclosure 2
Enclosure 3
F56 = 1
F75 = 1
2 Detector
4 Detector shield front
5 Detector Shield Rear
6 Cryostat Rear
7 Detector Rear

## Slide 8: Analytical Model
View factors in enclosure 1 sum to 1
Detector modelled as disk concentric with cryostat
  Circular detector with area equal to square detector (D = 117mm)
Shield and detector assumed perfectly coplanar (F24 = 0)
Shield perfectly floating (heat in = heat out)

## Slide 9: Analytical Model - Results
Heat flux to detector face | Q2 | 3.87 W
Shield temp | T4 | 256 K
Heat flux to shield (from encl 1) | Q4 | 0.32 W

## Slide 10: ANSYS Model - setup
Surface-surface radiation between cryostat, window, detector
Open enclosure, ambient temp = 256 K (shield temp)
Walls and window use shells
Detector uses volume elements
Detector shield not modelled
Small features and holes removed

## Slide 11: Boundary Conditions
Walls & window: 300K
  Window emissivity: 1
  Cryostat: 0.05 (shiny, polished)
Detector shield
  Modelled as “to ambient” radiation boundary @ 256 K
Detector and backplane
  Detector emissivity 1 (blackened)
  150 K cold strap applied
  G10 mounts ignored (perfect insulators)

## Slide 12: Mesh

## Slide 13: Surface-Surface Radiation
Enclosure type: open
Ambient temp: 256 K (temp of passive shield)
Emissivities:
  Window: 1
  Cryostat wall: 0.05
  Detector: 1

## Slide 14: Loads on detector
Radiation load to detector shield
  E = 0.05, T = 256 K
No radiation forward from backplane (occluded)
Cold strap: temp constraint @ 150 K

## Slide 15: Heat fluxes
 | Symbol | Analytical | ANSYS (ambient 300) | ANSYS (ambient 254)
Heat landing on detector face | Q2 | 3.95 | 3.87 | 3.87
Detector max temp | T2 | 150 | 162 | 162
Ambient temp set @ 300 K to validate analytical model
Set to 254 K for more accurate temp

## Slide 16: Detector temps

## Slide 17: Mechanical movement
Support constraints:
  Free to move radially per G10 flexures

## Slide 18: Radial Movement
Max: 0.13 mm

## Slide 19: Axial movement
Deviation across face 10 um
Most appears to be rigid-body movement

## Slide 20: Deviation from flatness
Get deviation from flat by subtracting out rigid-body motion
Range of deviation: +467/-428 nm

## Slide 21: Case 2: Remove ribs
Ribs removed by filling in pockets

## Slide 22: Temps with ribs removed
150.7 K max temp

## Slide 23: Displacement without ribs
Flatness deviation: +22/-11 nm

## Slide 24: Comparison of flatness change
Ribs removed
With ribs

## Slide 25: Flatness with ideal flexure mount
Temperature gradients kept from before
Ignore back plane
Constraint:
  Cylindrical CS, axial and tangential locked. Free in radial
  Simulates 3 ideal blade flexures

## Slide 26: Flatness with ideal flexure mounts
Deviation: +331/-339 nm

## Slide 27: Better active shielding
Simulated by:
  Set cryostat walls to 160K
  Set “to ambient” to 160 K
Analytical: heat flux to detector = 0.41 W
Ribs and mount kept from original
Qdet = 1.9 W
Analytical: 2.8 W
  Extra cold shield view factor likely the cause

## Slide 28: Better shielding effect on flatness
Deviation: +169/-162 nm

## Slide 29: Summary
Option | Min/Max face temp (K) | Max/min form deviation (nm)
Baseline – as designed | 151/162 | +467/-428
Remove ribs | 150.1 / 150.7 | +22/-11
Make ideal flexure mount (keep ribs) | 151/162 | +331/-339
Ideal active shielding (keep ribs) | 150.46 / 155.12 | +169/-162
Ribs removed and ideal shielding | 150.06 / 150.39 | 11.3/-6.7

## Slide 30: Conclusions
Removing ribs gives factor of 10 improvement in shape tolerance
  XY motion is dominated by absolute temperature. Z-axis is more what we see
Add active shield that covers all hot cryostat walls. Gives factor of 2.7 improvement for very little cost
Improving mount overconstraint has some effect, but minimal
In-plane movement of detector dominated by CTE. Almost all results show -21 um movement at edge
  Normal-to-beam movement will be Z-axis movement sin detector tilt angle

## Slide 31: Sensitivity to parameter changes
Question: how much should be concerned with maintaining strap temp, shield temp, detector current?
Method: bump parameters from equilibrium slightly, solve for image movement per unit change

## Slide 32: Parameters to be changed
Strap temp (simulates cooling power change)
Shield temp (varying +/- by 5K)
Heat load in detector face (3 W default, changed to 3.1 W)

## Slide 33: Uncorrectable XY movement
Subtract rigid-body movement and bulk thermal contraction
Represents movement not calibratable with single temp sensor
Subtract:
-Rigid Body Movement
-Uniform contraction
Linear regression
Residuals

## Slide 34: Uncorrectable XY movement
Single value for comparison
Use RMS of residuals

## Slide 35: Equibrium values
 | With ribs | Without ribs
Detector temp range (K) | 16.03464508 | 6.71352386
Max radial displacement (mm) | -0.01898411 | -0.02051212
RMS uncorrectable XY error (mm) | 0.00423649 | 0.00000472
Detector Flatness (mm) | 0.00065056 | 0.00027275

## Slide 36: Sensitivity Values
Sensitivity to | Unit | With Ribs | Without Ribs
Detector Power | mm/W | 5.92E-05 | 3.02E-05
Strap Temperature | mm/K | -1.43E-05 | -2.52E-06
Shield Temp | mm/K | 3.17E-06 | 2.15E-06
XY RMSE per unit change